## PFMEA
Potential
Failure Mode and Effects Analysis
System | [11 empty] | FMEA Number | P12345678
Component | [11 empty] | Prepared By | Class
Process |  |  |  |  | Design Responsibility |  |  | Class |  |  |  | FMEA Date (Orig.) | 2001-03-07
Model Year(s)/Project | 01 Hip Socket Replacement |  |  |  | Key Date |  |  | 2001-04-01 |  |  |  | FMEADate (Rev.) | NA
Core Team :
Class
Action Results
Design Item-Process Function/Design Function-Process Requirements* | Potential Failure Mode | Potential Effect(s) of Failure | Sev | Class | Potential Cause(s)/ Mechanism(s) of Failure | Occur | Prevention (Preventive Controls) | Current Process Design Controls (Detective Controls) | Detec | RPN | Recommended Action(s) | Responsibility & Target Completion Date | Actions Taken | Sev | Occ | Det | RPN
R1 :Obtain & verify Certs. | Misplaced / ommitted. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | 100% Insp.Weekly CFT review feedback. | 2 | 12 | No further action at this time.
Receiving Insp. Inspect raw materials as received (Sheet stock). | Thickness not checked. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
| Incorrect thickness conveyed. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
Re-package parts to send out. Route to dock. | Incorrect pack provided. | Components / Material damaged in transit. | 2 |  | Operator Error. Incorrect pack specified. | 3 | Operator Instructions with Visual Aids. | Quarterly Pack / Shipment Audits. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
ECT0 : Transport parts to EC1. | Delay in shipment. | Lost time Flow and delayed on time delivery. Customer dissatisfaction. | 2 |  | Transport / Plant carrier breakdown. | 3 | Scheduled PM. | Scheduled PM. | 2 | 12 | No further action at this time.
| Wrong parts shipped. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Barcode Traceability. Certs. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Pack damaged in shipment. | Rework / Delay | 2 |  | Operator Error. Incorrect pack specified. | 3 | Operator Instructions with Visual Aids. | Quarterly Pack / Shipment Audits.Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Cert lost in shipment. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. Incorrect pack specified. | 3 | Operator Instructions with Visual Aids. | Quarterly Pack / Shipment Audits.Weekly CFT review feedback. | 2 | 12 | No further action at this time.
EC1: Water jet (cut) parts to size. | Cut too large | Rework / Delay | 2 |  | Improper machine / tool setup. Inadequate Operator guidelines specified. Inadequate PM. | 3 | Operator Instructions with Visual Aids. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. Scheduled PM. | 2 | 12 | No further action at this time.
| Cut too small | Scrap (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / tool setup. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. Scheduled PM. | 2 | 18 | No further action at this time.
|  | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / tool setup. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. Scheduled PM. | 2 | 18 | No further action at this time.
| Not cut | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 2 | Operator Instructions with Visual Aids. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. Scheduled PM. | 2 | 8 | No further action at this time.
Route discrepant material to salvage for process internally / externally. | Not routed to salvage. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Red Tags at station. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
| Not re-processed internally. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Weekly CFT Reviews and feedback. | 2 | 24 | No further action at this time.
EC1: Package parts for shipment. Package certs. | Cert ommitted | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | 100 % Insp at R1. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Wrong cert provided | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | 100 % Insp at R1. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
| Incorrect parts packaged. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | 100 % Insp at R1. Weekly CFT review feedback. Barcode Traceability. | 2 | 24 | No further action at this time.
ECT1: Transport parts to Manufacturing. | Delay in shipment. | Delay | 2 |  | Transport / Plant carrier breakdown. | 3 | Scheduled PM. | Scheduled PM. | 2 | 12 | No further action at this time.
| Wrong parts shipped. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Insp. At R1. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Pack damaged in shipment. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. Incorrect pack specified. | 3 | Operator Instructions with Visual Aids. | Quarterly Pack / Shipment Audits.Weekly CFT review feedback. | 2 | 12 | No further action at this time.
|  | Scrap | 3 |  | Operator Error. Inadequate Operator guidelines specified. Incorrect pack specified. | 3 | Operator Instructions with Visual Aids. | Quarterly Pack / Shipment Audits.Weekly CFT review feedback. | 2 | 18 | No further action at this time.
| Cert lost in shipment. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Quarterly Pack / Shipment Audits. 100 % Insp at R1. Weekly CFT review feedback. Barcode Traceability. | 2 | 12 | No further action at this time.
RI : Obtain & verify Certs. | Verification ommitted. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | 100 % Insp. Barcode Traceability. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
| Incorrect verification transferred. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | 100 % Insp. Barcode Traceability. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
RI : Receiving Insp. Inspect raw materials as received (Thickness , size). | Thickness not checked. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
| Incorrect thickness conveyed. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
| Size not checked. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Incorrect size conveyed. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
Route discrepant material to salvage for process externally. | Not routed to salvage. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Red Tags at station. | 2 | 24 | No further action at this time.
Route material missing certs to salvage for re-process. | Not routed to salvage. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Weekly CFT Reviews and feedback. | 2 | 24 | No further action at this time.
RI: Re-package parts to send out for heat treat. Route to dock. | Incorrect pack. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. Incorrect pack specified. | 3 | Operator Instructions with Visual Aids. | Weekly CFT Reviews and feedback. | 2 | 12 | No further action at this time.
| Not routed to dock. Routed to line. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Weekly CFT Reviews and feedback.. Red Tags at station. | 2 | 24 | No further action at this time.
| Incorrect parts routed for transfer. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Weekly CFT Reviews and feedback. | 2 | 12 | No further action at this time.
ECT2 : Transport parts out for heat treat. | Delay in shipment. | Delay | 2 |  | Transport / Plant carrier breakdown. | 3 | Scheduled PM. | Scheduled PM. | 2 | 12 | No further action at this time.
| Wrong parts shipped. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Insp. At R1. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Pack damaged in shipment. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. Incorrect pack specified. | 3 | Operator Instructions with Visual Aids. | Quarterly Pack / Shipment Audits. 100% Insp. At R1. | 2 | 12 | No further action at this time.
| Cert lost in shipment. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. Incorrect pack specified. | 3 | Operator Instructions with Visual Aids. | Quarterly Pack / Shipment Audits. 100% Insp. At R1. | 2 | 12 | No further action at this time.
EC2A : Heat treat at 900F for 4 hrs. | Not treated / partially treated. | Scrap (Displaces. Hazard to end user). | 4 | CC | Improper machine / tool setup. Inadequate Operator guidelines specified. Inadequate PM. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. Scheduled PM. | 2 | 16 | No further action at this time.
| Incorrect treat. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Improper machine / tool setup. Inadequate Operator guidelines specified. Inadequate PM. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. Scheduled PM. | 2 | 24 | No further action at this time.
EC2A : Package parts for shipment. Package certs. | Cert ommitted | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | 100% Insp. At R1. Weekly CFT Reviews and feedback. | 2 | 12 | No further action at this time.
| Wrong cert provided | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | 100% Insp. At R1. Weekly CFT Reviews and feedback. | 2 | 24 | No further action at this time.
| Incorrect parts packaged. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | 100% Insp. At R1. Weekly CFT Reviews and feedback. | 2 | 24 | No further action at this time.
ECT2 : Transport parts to Manufacturing. | Delay in shipment. | Delay | 2 |  | Transport / Plant carrier breakdown. | 3 | Scheduled PM. | Scheduled PM. | 2 | 12 | No further action at this time.
| Wrong parts shipped. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Insp. At R1. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Pack damaged in shipment. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. Incorrect pack specified. | 3 | Operator Instructions with Visual Aids. | Quarterly Pack / Shipment Audits. 100% Insp. At R1. Weekly CFT Reviews and feedback. | 2 | 12 | No further action at this time.
| Cert lost in shipment. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. Incorrect pack specified. | 3 | Operator Instructions with Visual Aids. | Quarterly Pack / Shipment Audits. 100% Insp. At R1. Weekly CFT Reviews and feedback. | 2 | 12 | No further action at this time.
RI : Obtain & verify Certs. | Verification ommitted. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | 100% Insp. At R1. Weekly CFT Reviews and feedback. | 2 | 24 | No further action at this time.
| Incorrect verification transferred. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | 100% Insp. At R1. Weekly CFT Reviews and feedback. | 2 | 24 | No further action at this time.
RI : Receiving Insp. Inspect raw materials as received. | Inspection ommitted. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Weekly CFT Reviews and feedback. | 2 | 24 | No further action at this time.
Route material missing certs to salvage for hold. | Not routed to salvage. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Weekly CFT Reviews and feedback. | 2 | 24 | No further action at this time.
Route discrepant material to salvage for process externally. | Not routed to salvage. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Weekly CFT Reviews and feedback. Red Tags at station. | 2 | 24 | No further action at this time.
| Not re-processed internally. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Red Tags at station. Weekly CFT Review Feedback. | 2 | 24 | No further action at this time.
S1O1 : Polish : Radius wafers around edges to improve draw. | Radius ommitted. | Scrap (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | Yearly Operator Qualification. Weekly CFT review feedback. Scheduled PM. | 2 | 12 | No further action at this time.
|  | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | Yearly Operator Qualification. Weekly CFT review feedback. Scheduled PM. | 2 | 12 | No further action at this time.
| Excessive radius provided. | Scrap (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | Yearly Operator Qualification. Weekly CFT review feedback. Scheduled PM. | 2 | 18 | No further action at this time.
|  | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | Yearly Operator Qualification. Weekly CFT review feedback. Scheduled PM. | 2 | 18 | No further action at this time.
S2O1 : Draw sect. A2 to full height. Scribe lines at base to check for cracks. | Underdrawn. | Scrap (Difficult installation. Loss of function. A4 & A5 displace , Hazard to end user). | 4 | CC | Improper machine / Tool setup. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. Incorrect drawing compound (lube) specified / used.Inadequate PM. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | C/Np Chart / Lot.Yearly Operator Qualification. Weekly CFT review feedback. Scheduled PM. | 2 | 16 | Review feasability for multi-stage Prog. Die. | MW 4/24/01
| Overdrawn. | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | C/Np Chart / Lot. Yearly Operator Qualification. Weekly CFT review feedback.Scheduled PM. | 2 | 24 | Review feasability for multi-stage Prog. Die. | MW 4/24/01
|  | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | C/Np Chart / Lot. Yearly Operator Qualification. Weekly CFT review feedback.Scheduled PM. | 2 | 18 | Review feasability for multi-stage Prog. Die. | MW 4/24/01
S2O2 : Inspect for cracks | Not inspected for cracks. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | C/Np Chart / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
| Good material rejected. | Scrap | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
|  | Rework / Delay | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
Route discrepant material to salvage for process internally / externally. | Not routed to salvage. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Red Tags at station. Weekly CFT Review Feedback. | 2 | 24 | No further action at this time.
| Not re-processed internally. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Red Tags at station. Weekly CFT Review Feedback. | 2 | 24 | No further action at this time.
S3O1 : Inspect Dome thickness (3 pcs.) | Not inspected for thickness. Parts too thick. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
| Not inspected for thickness. Parts too thin. | Scrap (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
|  | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
Route discrepant material to salvage for process internally / externally. | Not routed to salvage. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Red Tags at station. Weekly CFT Review Feedback. | 2 | 24 | No further action at this time.
| Not re-processed internally. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Red Tags at station. Weekly CFT Review Feedback. | 2 | 24 | No further action at this time.
S3O2 : Center drill , drill , ream top dome hole. | Hole not drilled / partially drilled. | Rework / Delay. Flange holes drilled out of alignment further on in process.(Difficult installation. F1 displaces / galls. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing / tooling provided. Operator Error. Inadequate Operator guidelines specified. Incorrect cutting fluid specified / used. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | C / Np chart / lot.Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Hole drilled out of location. | Scrap. Flange holes drilled out of alignment further on in process.(Difficult installation. F1 displaces / galls. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing / tooling provided. Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
|  | Rework / Delay. Flange holes drilled out of alignment further on in process.(Difficult installation. F1 displaces / galls. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing / tooling provided. Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
| Hole drilled oversized. | Scrap. (Displaces. Hazard to end user). | 4 | CC | Improper machine / Tool setup. Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
|  | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Improper machine / Tool setup. Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
| Hole drilled undersized. | Rework / Delay. Flange holes drilled out of alignment further on in process.(Difficult installation. F1 displaces / galls. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing / tooling provided. Operator Error. Inadequate Operator guidelines specified. Incorrect cutting fluid specified / used. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Hole not reamed. | Rework / Delay. Flange holes drilled out of alignment further on in process.(Difficult installation. F1 displaces / galls. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Operator Error. Inadequate Operator guidelines specified. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | C / Np chart / lot.Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Hole reamed out of location. | Scrap. (Difficult installation. F1 displaces / galls. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing / tooling provided. Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
|  | Rework / Delay. Flange holes drilled out of alignment further on in process.(Difficult installation. F1 displaces / galls. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing / tooling provided. Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
S3O3 : Drill wire triangular holes in A2. | Hole(s) not drilled / partially drilled. | Rework / Delay. (Insufficient access. Non-optimal cement coverage and bonding potential. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing / tooling provided. Operator Error. Inadequate Operator guidelines specified. Incorrect cutting fluid specified / used. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | C / Np chart / lot.Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Hole(s) drilled out of location. | Scrap. (Insufficient access. Non-optimal cement coverage and bonding potential. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing / tooling provided. Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
|  | Rework / Delay. (Insufficient access. Non-optimal cement coverage and bonding potential. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing / tooling provided. Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
| Hole(s) drilled oversized. | Scrap. (Insufficient access. Non-optimal cement coverage and bonding potential. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing / tooling provided. Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
|  | Scrap. (Insufficient access. Non-optimal cement coverage and bonding potential. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing / tooling provided. Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
| Hole(s) drilled undersized. | Scrap. (Insufficient access. Non-optimal cement coverage and bonding potential. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing / tooling provided. Operator Error. Inadequate Operator guidelines specified. Incorrect cutting fluid specified / used. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
S3O4 : Deburr triangular holes in A2 for wire EDM. | Hole(s) not deburred / under de-burred. | Rework / Delay (Potential Blood Poisoning.Hazard to end user). | 4 | CC | Inadequate fixturing / tooling provided. Operator Error. Inadequate Operator guidelines specified. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | C / Np chart / lot.Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 16 | No further action at this time.
| Hole(s) over de-burred. | Scrap | 3 |  | Inadequate fixturing / tooling provided. Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | C / Np chart / lot.Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
|  | Rework / Delay | 2 |  | Inadequate fixturing / tooling provided. Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | C / Np chart / lot.Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
S3O5 : Draw. | Overdrawn. | Scrap. (Difficult installation. F1 displaces / galls. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | Review feasability for multi-stage Prog. Die. | MW 4/24/01
|  | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | Review feasability for multi-stage Prog. Die. | MW 4/24/01
| Underdrawn. | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. Incorrect drawing compound (lube) specified / used. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | Review feasability for multi-stage Prog. Die. | MW 4/24/01
S3O6 : Inspect drilled holes. | Drill hole inspection ommitted. | Rework / Delay. (Difficult installation. F1 displaces / galls. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
| Good material rejected. | Scrap | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
|  | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
Route discrepant material to salvage for process internally / externally. | Not routed to salvage. | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Red Tags at station. Weekly CFT Review Feedback. | 2 | 18 | No further action at this time.
| Not re-processed internally. | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Red Tags at station. Weekly CFT Review Feedback. | 2 | 18 | No further action at this time.
ECT3 : Transport parts to Manufacturing (EC2). | Delay in shipment. | Delay | 2 |  | Transport / Plant carrier breakdown. | 3 | Scheduled PM. | Scheduled PM. | 2 | 12 | No further action at this time.
| Wrong parts shipped. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Insp. At R1. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Pack damaged in shipment. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. Incorrect pack specified. | 3 | Operator Instructions with Visual Aids. | Quarterly Pack / Shipment Audits. 100% Insp. At R1. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Cert lost in shipment. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. Incorrect pack specified. | 3 | Operator Instructions with Visual Aids. | Quarterly Pack / Shipment Audits. 100% Insp. At R1. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
EC2B : Heat treat at 900F for 4 hrs. | Not treated / partially treated. | Scrap (Displaces. Hazard to end user). | 4 | CC | Improper machine / tool setup. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. Scheduled PM. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 16 | No further action at this time.
|  | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Improper machine / tool setup. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. Scheduled PM. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
| Incorrect treat. | Scrap (Displaces. Hazard to end user). | 4 | CC | Improper machine / tool setup. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. Scheduled PM. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
|  | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Improper machine / tool setup. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. Scheduled PM. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
EC2B : Package parts for shipment. Package certs. | Cert ommitted | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | 100% Insp. At R1. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Wrong cert provided | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | 100% Insp. At R1. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Incorrect parts packaged. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Insp. At R1. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
ECT4 : Transport parts to Manufacturing . | Delay in shipment. | Delay | 2 |  | Transport / Plant carrier breakdown. | 3 | Scheduled PM. | Scheduled PM. | 2 | 12 | No further action at this time.
| Wrong parts shipped. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Insp. At R1. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Pack damaged in shipment. | Scrap | 3 |  | Operator Error. Inadequate Operator guidelines specified. Incorrect pack specified. | 3 | Operator Instructions with Visual Aids. | Quarterly Pack / Shipment Audits. 100% Insp. At R1. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
|  | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. Incorrect pack specified. | 3 | Operator Instructions with Visual Aids. | Quarterly Pack / Shipment Audits. 100% Insp. At R1. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Cert lost in shipment. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. Incorrect pack specified. | 3 | Operator Instructions with Visual Aids. | Quarterly Pack / Shipment Audits. 100% Insp. At R1. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
RI : Obtain & verify Certs. | Verification ommitted. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Incorrect verification transferred. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
RI : Receiving Insp. Inspect raw materials as received. | Inspection ommitted. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
Route material missing certs to salvage for hold. | Not routed to salvage. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
Route discrepant material to salvage for process externally. | Not routed to salvage. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. Red tags at station. | 2 | 24 | No further action at this time.
S4O1 : Wire EDM. Wire cut profiles (A1,A2,A3) and triangular windows in A2 | Overcut on profiles. | Scrap. (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / tool setup. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. Improper cutting fluid specified / used. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. Scheduled Pm. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
|  | Rework / Delay. (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / tool setup. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. Improper cutting fluid specified / used. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. Scheduled PM. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
| Undercut on profiles. | Scrap. (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / tool setup. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. Improper cutting fluid specified / used. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. Scheduled PM. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
|  | Rework / Delay. (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / tool setup. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. Improper cutting fluid specified / used. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. Scheduled PM. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Profiles not cut. | Scrap. (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / tool setup. Operator Error. Inadequate Operator guidelines specified. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
|  | Rework / Delay. (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / tool setup. Operator Error. Inadequate Operator guidelines specified. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Overcut on windows. | Scrap. (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / tool setup. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. Improper cutting fluid specified / used. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. Scheduled PM. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
|  | Rework / Delay. (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / tool setup. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. Improper cutting fluid specified / used. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. Scheduled PM. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Undercut on windows. | Scrap. (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / tool setup. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. Improper cutting fluid specified / used. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. Scheduled PM. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
|  | Rework / Delay. (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / tool setup. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. Improper cutting fluid specified / used. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. Scheduled PM. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Windows not cut. | Scrap | 3 |  | Improper machine / tool setup. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. Improper cutting fluid specified / used. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. Scheduled PM. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
|  | Rework / Delay | 2 |  | Improper machine / tool setup. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. Improper cutting fluid specified / used. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. Scheduled PM. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 8 | No further action at this time.
S4O2 : Inspect parts to print. | Not inspected . | Rework / Delay. (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
| Good material rejected. | Scrap | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
|  | Rework / Delay | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
Route discrepant material to salvage for process internally / externally. | Not routed to salvage. | Rework / Delay. (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. Red tags at station. | 2 | 18 | No further action at this time.
| Not re-processed internally. | Rework / Delay. (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
S5O1 : Standard Mill. Mill side profiles | Side profiles over milled. | Scrap. (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / tool setup. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. Scheduled PM. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
|  | Rework / Delay. (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / tool setup. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. Scheduled PM. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
| Side profiles under milled. | Rework / Delay. (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / tool setup. Operator Error. Inadequate fixturing provided. Inadequate Operator guidelines specified. Inadequate PM. Improper cutting fluid specified / used. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. Scheduled PM. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Side profiles not milled. | Rework / Delay. (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
S5O2 : Deburr milled surfaces. | Surfaces not de- burred / under de-burred. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Inadequate tooling / fixturing provided. Operator Error. Inadequate Operator guidelines specified. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 16 | No further action at this time.
| Surfaces over de-burred. | Scrap | 3 |  | Inadequate tooling / fixturing provided. Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
|  | Rework / Delay | 3 |  | Inadequate tooling / fixturing provided. Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
S5O3 : Inspect for cracks. | Not inspected for cracks. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 2 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 16 | No further action at this time.
| Good material rejected. | Scrap | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 2 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
|  | Rework / Delay | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 2 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
Route discrepant material to salvage for process internally / externally. | Not routed to salvage. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. Red tags at station. | 2 | 24 | No further action at this time.
| Not re-processed internally. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
S6O1 : Small Press. Punch holes in flange A3 to final size. | Hole(s) not punched / partially punched. | Rework / Delay (Difficult installation. F1 displaces / galls. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. Incorrect drawing compound (lube) specified / used. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. Scheduled PM. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Hole(s) punched out of location. | Scrap. (Difficult installation. F1 displaces / galls. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / tool setup. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. Scheduled PM. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
|  | Rework / Delay (Difficult installation. F1 displaces / galls. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / tool setup. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. Scheduled PM. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
| Hole(s) punched oversized. | Scrap. (Difficult installation. F1 displaces / galls. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / tool setup. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. Scheduled PM. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
|  | Rework / Delay (Difficult installation. F1 displaces / galls. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / tool setup. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. Scheduled PM. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Hole(s) punched undersized. | Rework / Delay (Difficult installation. F1 displaces / galls. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. Incorrect drawing compound (lube) specified / used. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. Scheduled PM. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
S6O2 : Inspect hole diameters. | Hole(s) diameters not inspected . | Rework / Delay (Difficult installation. F1 displaces / galls. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
| Good material rejected. | Scrap | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
|  | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
Route discrepant material to salvage for process internally / externally. | Not routed to salvage. | Rework / Delay (Difficult installation. F1 displaces / galls. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. Red tags at station. | 2 | 18 | No further action at this time.
| Not re-processed internally. | Rework / Delay (Difficult installation. F1 displaces / galls. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
S6O3 : Inspect distance of holes to edge of part. | Hole(s) locations not inspected . | Rework / Delay (Difficult installation. F1 displaces / galls. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
| Good material rejected. | Scrap | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
|  | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
Route discrepant material to salvage for process internally / externally. | Not routed to salvage. | Rework / Delay (Difficult installation. F1 displaces / galls. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. Red tags at station. | 2 | 18 | No further action at this time.
| Not re-processed internally. | Rework / Delay (Difficult installation. F1 displaces / galls. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
S7O1 : Standard Mill. Countersink holes in flanges. | Hole(s) not countersunk / partially countersunk. | Rework / Delay (Difficult installation. F1 displaces / galls. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / tool setup. Operator Error. Inadequate fixturing provided. Inadequate Operator guidelines specified. Inadequate PM. Improper cutting fluid specified / used. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Hole(s) countersink out of location. | Scrap. (Difficult installation. F1 displaces / galls. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / tool setup. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
|  | Rework / Delay (Difficult installation. F1 displaces / galls. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / tool setup. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
| Hole(s) countersink oversized. | Scrap. (Difficult installation. F1 displaces / galls. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / tool setup. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
|  | Rework / Delay (Difficult installation. F1 displaces / galls. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / tool setup. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
S7O2 : Inspect countersinks (centered). | Not inspected for countersink centering. | Rework / Delay (Difficult installation. F1 displaces / galls. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
| Good material rejected. | Scrap | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
|  | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
Route discrepant material to salvage for process internally / externally. | Not routed to salvage. | Rework / Delay (Difficult installation. F1 displaces / galls. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. Red tags at station. | 2 | 18 | No further action at this time.
| Not re-processed internally. | Rework / Delay (Difficult installation. F1 displaces / galls. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
S7O3 : Inspect countersinks (Dia.). | Not inspected for countersink diameter. | Rework / Delay (Difficult installation. F1 displaces / galls. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
| Good material rejected. | Scrap | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
|  | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
Route discrepant material to salvage for process internally / externally. | Not routed to salvage. | Rework / Delay (Difficult installation. F1 displaces / galls. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. Red tags at station. | 2 | 18 | No further action at this time.
| Not re-processed internally. | Rework / Delay (Difficult installation. F1 displaces / galls. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
S8O1 : Polish. Break edges in grooves , windows and on wire cut edge. | Groove edges not broke / underbroke. | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Groove edges overbroke. | Scrap. (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
|  | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
| Window edges not broke / underbroke. | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Window edges overbroke. | Scrap. (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
|  | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
| Wire cut edges not broke / underbroke. | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Wire cut edges overbroke. | Scrap. (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
|  | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
S9O1 : Polish. Pre-form flange and hook. | Flange and hook not pre-formed / under formed. | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Flange and hook over formed. | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
S9O2 : Inspect material thickness in hook areas. | Hook area thickness not inspected /partial inspection.. | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
| Good material rejected. | Scrap | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
|  | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
Route discrepant material to salvage for process internally / externally. | Not routed to salvage. | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. Red tags at station. | 2 | 18 | No further action at this time.
| Not re-processed internally. | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
S10O1 : Small Press. Form flange and hook. | Flange and hook not formed / under formed. | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Flange and hook over formed. | Scrap. (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
|  | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
S10O2 : Inspect material thickness in hook areas. | Hook area thickness not inspected /partial inspection.. | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
| Good material rejected. | Scrap | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
|  | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
Route discrepant material to salvage for process internally / externally. | Not routed to salvage. | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. Red tags at station. | 2 | 18 | No further action at this time.
| Not re-processed internally. | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
S11O1 : Polish. Perform final bend (a1) on large flange A3. | Bend ommitted / underbent. | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Overbent. | Scrap. (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
|  | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
S11O2 : Inspect Thickness , bends and hole Dia. | Thickness not inspected /partial inspection.. | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
| Bend not inspected. | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
| Hole dia. Not inspected. | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
| Good material rejected. | Scrap | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
|  | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
Route discrepant material to salvage for process internally / externally. | Not routed to salvage. | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. Red tags at station. | 2 | 18 | No further action at this time.
| Not re-processed internally. | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
S12O1 : Polish. Trim hook end (A1) to length. Polish to length. | Hook end not trimmed to length / undertrimmed. | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Hook end overtrimmed. | Scrap (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
|  | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Hook end not polished to length. | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Hook end over polished. | Scrap (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
|  | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
S12O2 : Add full radius on end of hook (A1). | Radius ommitted / too small. | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Radius too large. | Scrap (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
|  | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
S13O1 : Standard Mill. Drill and countersink hole in A1. | Hole not drilled / partially drilled. | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. Inadequate PM. Improper cutting fluid specified / used. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Hole drilled out of location. | Scrap (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / tool setup. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
|  | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / tool setup. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
| Hole drilled oversized. | Scrap (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
|  | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
| Hole drilled undersized. | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Hole not countersunk / partially countersunk. | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / tool setup. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. Improper cutting fluid specified / used. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Hole countersink out of location. | Scrap (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
|  | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
| Hole countersink oversized. | Scrap (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
|  | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
S13O2 : Inspect hole in A1. | Hole dia. Not inspected. | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
| Good material rejected. | Scrap | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
|  | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
Route discrepant material to salvage for process internally / externally. | Not routed to salvage. | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. Red tags at station. | 2 | 18 | No further action at this time.
| Not re-processed internally. | Rework / Delay (Difficult installation. Loss of function. Extreme physical discomfort and customer dissatisfaction). | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
S14O1: Standard Mill. Drill all dome holes in A2. | Hole(s) not drilled / partially drilled. | Rework / Delay (Difficult installation. F1 displaces / galls. A2 displaces.Hazard to end user). | 4 | CC | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. Inadequate PM. Improper cutting fluid specified / used. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 16 | No further action at this time.
| Hole drilled out of location. | Scrap | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. Inadequate PM. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
|  | Rework / Delay (Difficult installation. F1 displaces / galls. A2 displaces.Hazard to end user). | 4 | CC | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. Inadequate PM. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
| Hole drilled oversized. | Scrap (Difficult installation. F1 displaces / galls. A2 displaces.Hazard to end user). | 4 | CC | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
|  | Rework / Delay (Difficult installation. F1 displaces / galls. A2 displaces.Hazard to end user). | 4 | CC | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
| Hole drilled undersized. | Rework / Delay (Difficult installation. F1 displaces / galls. A2 displaces.Hazard to end user). | 4 | CC | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. Improper cutting fluid specified / used. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 16 | No further action at this time.
S14O2 : Inspect Dome holes in A2. | Hole(s) dia. Not inspected. | Rework / Delay (Difficult installation. F1 displaces / galls. A2 displaces.Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
| Good material rejected. | Scrap | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
|  | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
Route discrepant material to salvage for process internally / externally. | Not routed to salvage. | Rework / Delay (Difficult installation. F1 displaces / galls. A2 displaces.Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. Red tags at station. | 2 | 24 | No further action at this time.
| Not re-processed internally. | Rework / Delay (Difficult installation. F1 displaces / galls. A2 displaces.Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
S15O1 : Standard Mill. Countersink dome holes in A2. | Hole(s) not countersunk / partially countersunk. | Rework / Delay (Difficult installation. F1 displaces / galls. A2 displaces.Hazard to end user). | 4 | CC | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. Inadequate PM. Improper cutting fluid specified / used. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 16 | No further action at this time.
| Hole(s) countersink out of location. | Scrap (Difficult installation. F1 displaces / galls. A2 displaces.Hazard to end user). | 4 | CC | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. Inadequate PM. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
|  | Rework / Delay (Difficult installation. F1 displaces / galls. A2 displaces.Hazard to end user). | 4 | CC | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. Inadequate PM. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
| Hole(s) countersink oversized. | Scrap (Difficult installation. F1 displaces / galls. A2 displaces.Hazard to end user). | 4 | CC | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
|  | Rework / Delay (Difficult installation. F1 displaces / galls. A2 displaces.Hazard to end user). | 4 | CC | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
S15O2 : Inspect Dome holes in A2. | Hole(s) dia. Not inspected. | Rework / Delay (Difficult installation. F1 displaces / galls. A2 displaces.Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
| Good material rejected. | Scrap | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
|  | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
Route discrepant material to salvage for process internally / externally. | Not routed to salvage. | Rework / Delay (Difficult installation. F1 displaces / galls. A2 displaces.Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. Red tags at station. | 2 | 24 | No further action at this time.
| Not re-processed internally. | Rework / Delay (Difficult installation. F1 displaces / galls. A2 displaces.Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
S16O1 : Polish. Break edges. | Edges not polished / partially polished. | Rework / Delay. (Potential Blood Poisoning.Hazard to end user). | 4 | CC | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 16 | No further action at this time.
| Edges over polished. | Scrap | 2 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
|  | Rework / Delay | 2 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
S17O1 : Tumble Finish. Gravi-flo parts in brown media for 1 Hr. | Not tumbled. | Rework / Delay. (Potential Blood Poisoning.Hazard to end user). | 4 | CC | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. Inadequate PM. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 16 | No further action at this time.
| Not tumbled for correct period. | Rework / Delay. (Potential Blood Poisoning.Hazard to end user). | 4 | CC | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. Inadequate PM. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
R1 : Re-package parts to send out. Route to dock. | Not packaged correctly. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Wrong parts packed. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Weekly CFT review feedback. | 2 | 24 | No further action at this time.
| Not routed to dock for transfer. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Incorrect parts routed for transfer. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Weekly CFT review feedback. | 2 | 24 | No further action at this time.
ECT5 : Transport parts out for Medcast check (Zyglo). | Delay in shipment. | Delay | 2 |  | Transport / Plant carrier breakdown. | 3 | Scheduled PM. | Scheduled PM. | 2 | 12 | No further action at this time.
| Wrong parts shipped. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Weekly CFT review feedback. | 2 | 24 | No further action at this time.
| Pack damaged in shipment. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. Incorrect pack specified. | 3 | Operator Instructions with Visual Aids. | Quarterly Pack / Shipment Audits. 100% Insp. At Medcast. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Cert lost in shipment. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. Incorrect pack specified. | 3 | Operator Instructions with Visual Aids. | Quarterly Pack / Shipment Audits. 100% Insp. At Medcast. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
EC2 : Inspect parts for cracks (Zyglo Dye Penetrant test). | Not inspected for cracks. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. Inadequate PM. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | C / Np chart / lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 16 | No further action at this time.
| Good material rejected. | Scrap | 3 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. Inadequate PM. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
|  | Rework / Delay | 2 |  | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. Inadequate PM. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 8 | No further action at this time.
Route discrepant material to salvage for process internally / externally. | Not routed to salvage. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. Red tags at station. | 2 | 24 | No further action at this time.
EC2 : Package parts and certs. For shipment. | Cert ommitted | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | 100% Insp. At R1. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Wrong cert provided | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | 100% Insp. At R1. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Incorrect parts packaged. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | 100% Insp. At R1. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
ECT6 : Transport parts to Manufacturing. | Delay in shipment. | Delay | 2 |  | Transport / Plant carrier breakdown. | 3 | Scheduled PM. | Scheduled PM. | 2 | 12 | No further action at this time.
| Wrong parts shipped. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | 100% Insp. At R1. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Pack damaged in shipment. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. Incorrect pack specified. | 3 | Operator Instructions with Visual Aids. | Quarterly Pack / Shipment Audits. 100% Insp. At R1. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Cert lost in shipment. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. Incorrect pack specified. | 3 | Operator Instructions with Visual Aids. | Quarterly Pack / Shipment Audits. 100% Insp. At R1. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
RI : Obtain & verify Certs. | Verification ommitted. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | 100% Insp. At R1. Weekly CFT review feedback. Barcode traceability. | 2 | 24 | No further action at this time.
| Incorrect verification transferred. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | 100% Insp. At R1. Weekly CFT review feedback. Barcode traceability. | 2 | 12 | No further action at this time.
Route material missing certs to salvage for process internally / externally. | Not routed to salvage. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. Red tags at station. | 2 | 24 | No further action at this time.
RI : Receiving Insp. Check for cracks. | Not inspected for cracks. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
| Good material rejected. | Scrap | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
|  | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
Route discrepant material to salvage for process internally / externally. | Not routed to salvage. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. Red tags at station. | 2 | 24 | No further action at this time.
| Not processed internally. | Discrepant material introduced to line. Scrap / Delay. (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
S18O1 : Blast. Cosmetic blast (30 grit ALOX). | Not blasted . | Rework / Delay. (Potential Blood Poisoning.Hazard to end user). | 4 | CC | Improper machine / Tool setup. Inadequate fixturing provided. Operator Error.Inadequate Operator guidelines specified. Inadequate PM. | 2 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 16 | No further action at this time.
| Blasted with incorrect grit. | Rework / Delay. (Potential Blood Poisoning.Hazard to end user). | 4 | CC | Improper machine / Tool setup. Operator Error. Inadequate Operator guidelines specified. Inadequate PM. | 3 | Operator Instructions with Visual Aids. Operator training and qualification on equipment. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
S19O1 : Compile and review Master Record packet. | Not compiled. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
| Incorrect packet compiled. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
| Not reviewed / insufficiently reviewed. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
S19O2 : Inspect for discoloration. Re-blast as needed. | Not inspected for dis-coloration. | Rework / Delay. (Potential Blood Poisoning.Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 2 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 16 | No further action at this time.
| Not re-blasted when required. | Rework / Delay. (Potential Blood Poisoning.Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 2 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 16 | No further action at this time.
| Re-blasted when not required. | Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
S19O3 : Inspect thickness in hook areas. | Hook area thickness not inspected /partial inspection.. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
| Good material rejected. | Scrap | 3 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 18 | No further action at this time.
|  | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 12 | No further action at this time.
Route discrepant material to salvage for process internally / externally. | Not routed to salvage. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. Red tags at station. | 2 | 24 | No further action at this time.
| Not processed internally. | Discrepant material introduced to line. Scrap / Delay. (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
S19O4 : Inspect radius. | Radius not inspected. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. Operator training on inspection equipment. Calibrated equipment. Specified Cal. Interval. | SPC / Lot. Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
Route discrepant material to salvage for process internally / externally. | Not routed to salvage. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
| Not re-processed internally. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
S19O5 : Inspect edges. | Edges not inspected. | Rework / Delay. (Potential Blood Poisoning.Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
Route discrepant material to salvage for process internally / externally. | Not routed to salvage. | Rework / Delay. (Potential Blood Poisoning.Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. Red tags at station. | 2 | 24 | No further action at this time.
| Not re-processed internally. | Rework / Delay. (Potential Blood Poisoning.Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Yearly Operator Qualification. Weekly CFT review feedback. | 2 | 24 | No further action at this time.
S20O1 : Clean , Pack and Ship. Clean parts. | Parts not cleaned / partially cleaned. | Rework / Delay. (Potential Blood Poisoning.Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. Incorrect cleaning materials specified / used. Excessive contamination in area. | 2 | Operator Instructions with Visual Aids. | Weekly CFT reviews and feedback. | 2 | 16 | Class 5 clean room. | MW 10/14/01
Reprocess parts (clean) if contaminated. | Not re-processed internally. | Rework / Delay. (Potential Blood Poisoning.Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 2 | Operator Instructions with Visual Aids. | Weekly CFT reviews and feedback. | 2 | 16 | No further action at this time.
S20O2 : Pack parts in bubble bags. | Part(s) not packed in bubble bag. | Contamination of parts. (Potential Blood Poisoning.Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Weekly CFT reviews and feedback. | 2 | 24 | No further action at this time.
S20O3 : Pack bags in box. | Wrong part(s) packed in box. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Weekly CFT reviews and feedback. | 2 | 24 | No further action at this time.
| Box underpacked. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Weekly CFT reviews and feedback. | 2 | 12 | No further action at this time.
| Box over packed. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Weekly CFT reviews and feedback. | 2 | 12 | No further action at this time.
S20O4 : Identify box (Label / Part number). | Label ommitted. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Weekly CFT reviews and feedback. | 2 | 12 | No further action at this time.
| Incorrect label installed. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Weekly CFT reviews and feedback. | 2 | 24 | No further action at this time.
| Label damaged / contaminated. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. Incorrect pack specified. | 3 | Operator Instructions with Visual Aids. | Quarterly Pack / Shipment Audits. 100% Insp. Biomet. Weekly CFT reviews and feedback. | 2 | 12 | No further action at this time.
S20O5 : Put heat code on V of C. | Heat code ommitted. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Weekly CFT reviews and feedback. | 2 | 24 | No further action at this time.
S20O6 : Audit device Master record (SOP 3027). | Not audited / partially audited. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Weekly CFT reviews and feedback. | 2 | 12 | No further action at this time.
R1 : Route parts to dock to send out. | Not routed to dock for transfer. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Weekly CFT reviews and feedback. | 2 | 12 | No further action at this time.
| Incorrect parts routed for transfer. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Weekly CFT reviews and feedback. | 2 | 12 | No further action at this time.
ECT7 : Transport parts out for Biomet for finish. | Delay in shipment. | Delay | 2 |  | Transport / Plant carrier breakdown. | 3 | Scheduled PM. | Scheduled PM. | 2 | 12 | No further action at this time.
| Wrong parts shipped. | Rework / Delay (Displaces. Hazard to end user). | 4 | CC | Operator Error. Inadequate Operator guidelines specified. | 3 | Operator Instructions with Visual Aids. | Weekly CFT reviews and feedback. | 2 | 24 | No further action at this time.
| Pack damaged in shipment. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. Incorrect pack specified. | 3 | Operator Instructions with Visual Aids. | Quarterly Pack / Shipment Audits. 100% Insp. Biomet. Weekly CFT reviews and feedback. | 2 | 12 | No further action at this time.
| Cert lost in shipment. | Rework / Delay | 2 |  | Operator Error. Inadequate Operator guidelines specified. Incorrect pack specified. | 3 | Operator Instructions with Visual Aids. | Quarterly Pack / Shipment Audits. 100% Insp. Biomet. Weekly CFT reviews and feedback. | 2 | 12 | No further action at this time.
PFMEA Statistics |  |  |  |  | 0<rpn<9 | 3 | 0<rpn<9 | 3 |  | AIAG | "O"Conversion
Characteristics Count |  | Average |  |  | rpn<19 | 240 | 10<rpn<19 | 237 |  | 6.7e-07 | 0
# CC's | 100 | Causes |  |  | rpn<29 | 324 | 20<rpn<29 | 84 |  | 6.67e-06 | 0.00042021
# SC's | 0 | Total |  |  | rpn<39 | 324 | 30<rpn<39 | 0 |  | 6.667e-05 | 0.017401
Total : | 100 | 1292 |  |  | rpn<49 | 324 | 40<rpn<49 | 0 |  | 0.0005 | 0
Sum of modes considered : 323 |  |  |  |  | rpn<59 | 324 | 50<rpn<59 | 0 |  | 0.0025 | 0
Index / Criteria | # |  |  |  | rpn<69 | 324 | 60<rpn<69 | 0 |  | 0.0125 | 0
PO=1 | 0 |  |  |  | rpn<79 | 324 | 70<rpn<79 | 0 |  | 0.05 | 0
PO=2 | 63 |  |  |  | rpn<89 | 324 | 80<rpn<89 | 0 |  | 0.125 | 0
PO=3 | 261 |  |  |  | rpn<99 | 324 | 90<rpn<99 | 0 |  | 0.33333 | 0
PO=4 | 0 |  |  |  | rpn<125 | 334 | 100<rpn<125 | 0 |  | 0.5 | 0
PO=5 | 0 |  |  |  | rpn<250 | 324 | 126<rpn<250 | 0
PO=6 | 0 |  |  |  | rpn<299 | 324 | 251<rpn<299 | 0 |  | Sum "U" | 0.017821 | Claims / 1000 | Less Recv. Insp.& Transport
PO=7 | 0 |  |  |  | rpn>299 | 0 | rpn>299 | 0 |  |  | Reliability | 17.8211 | 11.2074
PO=8 | 0 | [9 empty] | 0.98218
PO=9 | 0
PO=10 | 0
S=1 | 0
S=2 | 79
S=3 | 145
S=4 | 100
D=1 | 0
D=2 | 324
D=3 | 0
D=4 | 0
D=5 | 0
D=6 | 0
D=7 | 0
D=8 | 0
D=9 | 0
D=10 | 0

## Control Plan
CONTROL PLAN
| Prototype | Pre-launch |  | Production | Key Contact/Phone |  |  | Date (Orig.) |  | Date (Rev.)
Control Plan Number | [7 empty] | 2002-10-22
Part Number/Latest Change Level |  |  |  |  | Core Team |  |  | Customer Engineering Approval/Date (if req'd)
12345678 |  |  |  |  | Class
Part Name/Description |  |  |  |  | Supplier/Plant Approval/Date |  |  | Customer Quality Approval/Date (if req'd)
Hip Socket Replacement
Supplier/Plant |  |  | Supplier Code |  | Other Approval/Date (if req'd) |  |  | Other Approval/Date (if req'd)
IUSB
Special Characteristics |  |  |  | CC | Critical (Sev = 4) |  |  | SC | Significant ( Sev < 4 & RPN >/= 60)
Part Feature / | Process Name/ | Machine, Device, | Characteristics |  |  | Special | Methods
Process | Operation | Jig, Tools | No. | Product | Process | Char. | Product/Process | Evaluation/ | Sample |  | Control | Reaction Plan
Number | Description | for Mfg. |  |  |  | Class. | Specification/Tolerance | Measurement Technique | Size | Freq. | Method
A2 : S201 | OD Diameter of socket receptacle;R6 Inside Radius of socket receptacle : Draw sect. To full height. | 123456 : Micrometer ; XXYTZ5 : Fixture ; 99999876 : Height Gage. | 1 | X | X | CC | Ref dwg. 12345678 | OD Micrometer . (4) spot meas. About periphery , Subtract min. & max. and plot ovality. Measure height to assess radius. | 100% Insp. | Shift | X bar & R chart. Yearly operator qualification. Red tag process. | Containment , Rework / Scrap. Weekly CFT reviews and feedback.
A2 : S302 , S306 , S1401 , S1501 | Socket Dome hole :OD-Diameter of hole ;ID-Diameter of hole | Pin gages : 29292989 & 39393989 | 2 | X |  | CC | Ref dwg. 12345678 | Min & Max Pin gage on hole diameters. Go / No Go. | 100% Insp. | Shift | In process Checksheet. Yearly operator qualification.Red tag process. | Containment , Rework small / Scrap large. Weekly CFT reviews and feedback.
A2 : S306 , S1402 , S1502 | S2 adjacent hole to hole centerline spacing within and across rows ; | Plug gage : AXCD567 | 3 | X |  | CC | Ref dwg. 12345678 | Go / No Go. | 100% Insp. | Shift | In process Checksheet. Yearly operator qualification.Red tag process. | Containment , Rework small / Scrap large.Weekly CFT reviews and feedback.
A2 | D0 Hole shank depth from outside socket face to bottom | XXYTZ5 : Fixture ; 99999876 : Height Gage. | 4 | X |  | CC | Ref dwg. 12345678 | Dimensional | 5 | Hourly | X bar & R chart. Yearly operator qualification.Red tag process. | Containment , Rework small / Scrap large.Weekly CFT reviews and feedback.
R9 : S1904 | Radii | Radius gage : 66667898 | 5 | X |  | CC | Ref dwg. 12345678 | Go / No Go. | 100% Insp. | Shift | In process Checksheet. Yearly operator qualification.Red tag process. | Containment , Rework small / Scrap large / small.Weekly CFT reviews and feedback.
M : R1 , EC2A , EC2B | Materials : Hardness | AKATZ5 : Hardness Tester . | 6 | X |  | CC | Ref dwg. 12345678 | Hardness Test | 5 | Shift | R chart. Incoming and in process lot certifications. Yearly operator qualification.Red tag process.Barcode traceability. | Containment , Rework .Weekly CFT reviews and feedback.
S 202 , S503 , R1 | Materials : Integrity , Inspect for cracks. | NA | 7 |  | X | CC | Ref dwg. 12345678 | Visual | 100% Insp. | Shift | Visual , In process Checksheet. Yearly operator qualification.Red tag process. | Containment , Scrap .Weekly CFT reviews and feedback.
EC2 |  | External Supplier . | 7 |  | X | CC | Ref dwg. 12345678 | Zyglo penetrant tester. | 100% Insp. | Shift | Visual , In process Checksheet. Yearly operator qualification.Red tag process. | Containment , Scrap .Weekly CFT reviews and feedback.
P : S304 , S502 , S1601 ,S1701, S1801 , S1902 , S1904 , S1905 , ECT7 | Coating / Finish : Polish , break edges ; Tumble finish , Gravi-flo parts ; Cosmetic Blast ; Inspect for discoloration | ABCTZ5 : Profilometer | 8 | X | X | CC | Ref dwg. 12345678 | Profilometer. | 100% Insp. | Shift | Visual , In process Checksheet. Yearly operator qualification. Red tag process. | Containment , Rework .Weekly CFT reviews and feedback.
S2001 , S2002 | Hypoallergenic : Clean parts | WWWW112 : Analyzer ; V8YU324 : Dessicator ; KK3241Q : Scale | 9 | X | X | CC | Ref dwg. 12345678 | Chemical Analyzer and dessicator. | 5 | Shift | Incoming and in process contaminant check. 100% wash and dry prior to pack. Yearly operator qualification.Red tag process. Dock Audit. | Containment , Rework / Rewash .Weekly CFT reviews and feedback.
| Contaminant free. |  | 9 | X | X | CC
T1 : R1 , S301 , S1903 | Thickness of A1 , A2 & A3 | 9987678 Digital Calipers | 10 | X | X | CC | Ref dwg. 12345678 | Dimensional | 100% Insp. | Shift | X bar & R chart. Yearly operator qualification.Red tag process. Barcode traceability. | Containment , Scrap thin.Weekly CFT reviews and feedback.
R1 : EC2A , EC2B | Verify certifications.Route missing certs to salvage for hold. | NA | 11 |  | X | CC | Process Instruction #120 | Visual at each location. | 100% Insp. | Shift | Operator label initial. Yearly operator qualification.Red tag process. Barcode traceability. | Containment , Return / Rework .Weekly CFT reviews and feedback.
R1 , S202 , S301 , S503 , S1402 , S1502 , EC2 , S1903 , S1904 , S1905 | Route discrepant material to salvage for process internally / externally. | NA | 12 |  | X | CC | Process Instruction #121 | Visual at each location. | 100% Insp. | Shift | Red tag process. In process Checksheet. Yearly operator qualification.Red tag process. | Weekly CFT reviews and feedback.
EC1 , EC2A , EC2B , R1 , EC2 | Package parts and certifications for shipment. | NA | 13 |  | X | CC | Process Instruction #122 | Visual at each location. | Batch | Shift | Audit. Barcode traceability.Red tag process. | Weekly CFT reviews and feedback.
S2002 , S2004 , S2005 | Final package for shipment. Material segregation and package labeling. Place heat code. | NA | 14 |  | X | CC | Process Instruction #123 | Visual at each location. | 100% Insp. | Shift | Operator label initial. Yearly operator qualification.Red tag process. Barcode traceability. | Containment , Return / Rework .Weekly CFT reviews and feedback.
S1901 | Compile and review Master record packet. Final audit. | NA | 15 |  | X | CC | Process Instruction #124 | Dock Audit. Visual at dock location. | 100% Insp. | Shift | Operator label initial. Yearly operator qualification.Red tag process. Barcode traceability. | Containment , Return / Rework .Weekly CFT reviews and feedback.